## Doppler
COLOR FOETAL DOPPLER UNIT
S.N | Technical Specification | Bidder's Compliance Sheet
| COLOR FOETAL DOPPLER UNIT [Koshi Hospital,F/Y-2081/082] | Yes/No | Page no in Catalogue | Remarks
Manufacturer:
Brand:
Type/Model:
Country of Origin:
1.1 | Doppler foetal heart detector is a handheld ultrasound transducer used to detect the heartbeat of a foetus for prenatal care.
2 | Operational Requirements
2.1 | Shall be handheld, lightweight and easy to carry.
3 | System Configuration
3.1 | Doppler, Foetal Heart Detector, complete with accessories.
4 | Technical Specifications
4.1 | Doppler based foetal heart rate detector with built in speaker and colour LCD display
4.2 | Ultrasonic working frequency, approx.: 3MHz
4.3 | FHR measurement ranger and display range shall be approx. 50 - 210 bpm
4.4 | Doppler shall have countinous mode of operation
4.5 | Accuracy should be ±2 bpm
4.6 | Colour LCD/TFT display show real FHR value, average FHR value, battery status and volume level,etc.
4.7 | System shall report operational status, malfunctions and low battery with audio-visual alerts.
4.8 | Built-in loudspeaker with volume adjustment.
4.9 | The fetal doppler is designed with an overall sensitivity of approx ≥90dB, ensuring accurate and reliable detection of fetal heartbeats.
4.1 | Shall operate on rechargeable battery with original battery that provides a full charge, allowing atleast 3 hours of continuous use.
5 | ACCESSORIES
5.1 | Ultrasound gel - 2 tubes
5.2 | Carry bag - 1 set
5.3 | Rechargeable Battery - 2 set
5.4 | Battery charger - 1 set
5.3 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders Shall specify the quantity of every item included in their offer. (including items not listed above.)
6 | OPERATING ENVIRONMENT
6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power supply, climate, temperature, humidity, etc.
6.2 | Power Supply: 220 - 240 VAC, 50 Hz Single phased fitted with appropriate 3 pin plug (FLAT) or DC supply with adapter . The power cable Shall be at least 3m long.
7 | STANDARDS AND SAFETY REQUIREMENTS
7.1 | Shall submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
8 | TRAINING
8.1 | Shall provide user training (including application: how to use and maintain the equipment) to concerned user until complete familiarity with the system.
9 | WARRANTY
9.1 | Comprehensive warranty for 1 years on system.
10 | MAINTENANCE DURING SERVICE PERIOD
10.1 | During warranty period supplier Shall ensure corrective/breakdown maintenance whenever required.
12 | INSTALLATION, INSPECTION, COMMISSIONING
12.1 | The bidder Shall arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail
13 | DOCUMENTATION
13.1 | User (operating) manual in English .
13.2 | Service/ Maintenance manual in English .
13.4 | Manufacturer authorization letter

## CTG or NST
Technical Specification of NST/Cardiotocograph(CTG)
S.N | Technical Specification | Bidder's Compliance Sheet
| NST/Cardiotocograph(CTG) [Koshi Hospital,F/Y-2081/082] | Yes/No | Page no in Ref. Docs. | Remarks
Manufacturer:
Brand:
Type/Model:
Country of Origin:
1 | Description of Function
1.1 | CTG monitor is used in health facilities for routine assessment of foetal heart rate(FHR) and uterine contraction (UC), during labour and throughout delivery.
2 | Operational Requirements
2.1 | External foetal monitoring system, shall work on mains electric supply.
3 | System Configuration
3.1 | Cardiotocograph(CTG) monitor with complete accessories.+
4 | Technical Specifications
4.1 | It shall have ultrasonic transducer to measure foetal heart rate (FHR).
4.2 | Facility of calculation of FHR base line, variability, accelerations and decelerations.
4.3 | Facuility of audio-visual alert on loss of signals.
4.4 | Shall have pressure-sensitive transducer records uterine contractions (UC)
4.5 | Monitor fit with remote switch for event marking.
4.6 | Monitor should have provision to be used for twins.
4.7 | System should automatically recognizes specific transducers.
4.8 | Provision of performing self-test when device is turned on each time should be available.
4.9 | Large alphanumeric display provided with: FHR 1, FHR 2 and UC must be available.
4.1 | FHR range: approx. 50 to 240 bpm.
4.11 | UC range: approx. 0 to 100 units.
4.12 | Sustem should have reporting provision (operational status, malfunction transducers, out-of-paper) along with audio-visual alert.
4.13 | Printer:
a. | Built in high-resolution thermal printer, paper width approx. 150mm.
b. | Automatic and manual print-out mode.
c. | Prints FHR 1, FHR 2 and UC, displayed parameters and marked events.
d. | Printer resolution, approximately 1 bpm (FHR) and 1 unit UC.
e. | Paper speed adjustable: 1, 2 or 3 cm/min.
5 | ACCESSORIES
Shall come with:
5.1 | 1 x contraction transducer.
5.2 | 2 x Foetal heart Ultrasonic sensor.
5.3 | 1 x Remote switch event marker with cable.
5.4 | 2 x Foetal heart Ultrasonic sensor.
2 x Transducer belt 5 X 150 cm, length adjustable.
30 x thermal recording z-folded sheets.
2 x Bottle of ultrasound gel, approximately 250ml.
1 x set of fuse.
1 x Plastic protective dust-cover.
5.3 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders Shall specify the quantity of every item included in their offer. (including items not listed above.)
6 | OPERATING ENVIRONMENT
6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power supply, climate, temperature, humidity, etc.
6.2 | Power Supply: 220 - 240 VAC, 50 Hz Single phased fitted with appropriate 3 pin plug (FLAT) or DC supply with adapter . The power cable Shall be at least 3m long.
7 | STANDARDS AND SAFETY REQUIREMENTS
7.1 | Shall submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
8 | TRAINING
8.1 | Shall provide user training (including application: how to use and maintain the equipment) to concerned user until complete familiarity with the system.
9 | WARRANTY
9.1 | Comprehensive warranty for 2 years on system.
10 | MAINTENANCE DURING SERVICE PERIOD
10.1 | During warranty period supplier Shall ensure corrective/breakdown maintenance whenever required.
12 | INSTALLATION, INSPECTION, COMMISSIONING
12.1 | The bidder Shall arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail
13 | DOCUMENTATION
13.1 | User (operating) manual in English .
13.2 | Service/ Maintenance manual in English .
13.4 | Manufacturer authorization letter